## Elect Test Limits
Electrical Tests per AS/NZS 3551 |  |  |  | Rev: Nov 2019
Test | Class | Limit | Limit | Connection | Notes:
Protective Earth | 1 | < 0.2 ohms* |  | E - earthed parts | 1
Insulation Resistance | 1,2 | > 1 Meg ohm* |  | P/N - E | 2
P/N - Encl
P/N - AP
Touch Current (Enclosure leakage): Normal | 1 | < 100uA |  | E - Encl
Touch Current: Neut Open | 1 | < 500uA |  | E - Encl
Touch Current: Earth Open | 1 | < 500uA |  | E - Encl
Earth Leakage Current: Normal | 1 | < 5mA |  | E
Earth Leakage Current: Neut Open | 1 | < 10mA |  | E
|  | B,BF | CF
Patient Leakage Current: Normal | 1,2 | < 100uA | <10uA | each AP type
Patient Leakage Current: Neut Open | 1,2 | < 500uA | < 50uA | each AP type
Patient Leakage Current: Earth Open | 1,2 | < 500uA | < 50uA | each AP type
|  | BF | CF
Mains on applied part (BF) | 1,2 | < 5mA | < 50uA | each AP type
Note 1: 0.1+0.2=0.3 if detachable lead
Note 2: 1 to 10 conditionally permitted. Unconditional if > 10M
This is a simplified summary of electrical tests. Physical inspection and non electrical requirements also apply;
eg equipment labelling, detachable cords "tagged" etc
Electrical Tests per AS/NZS 3760
Test | Class | Limit | Limit | Connection | Notes:
Protective Earth | 1 | < 1 ohm | fix-wired<2 | E - earthed parts | 5
Insulation Resistance ** | 1,2 | > 1 Meg ohm |  | P/N - E | 3
Earth Leakage Current: Normal ** | 1 | < 5mA |  | E | 3
Touch Current (Enclosure leakage): Normal | 2 | < 1mA |  | E - Encl | 4
Note 3: AS/NZS 3760 requires ** EITHER Insulation OR Leakage Current test (not necessarily both)
Leakage can be measured Directly, with Clamp Meter, or by Differential method.
Note 4: AS/NZS 3760 calls this Class II Leakage Current |  |  | Note5: fix-wired is Clin Eng special case
Exceptions:
IEC 60601-2-54:2009 Mobile X-Ray
Touch Current (Enclosure leakage): Normal | 1 | < 100uA
Touch Current: Neut Open | 1 | < 2mA
Touch Current: Earth Open | 1 | < 2mA
Earth Leakage Current: Normal | 1 | < 2.5mA
Earth Leakage Current: Neut Open | 1 | < 5mA
no AP Tests